Optimal settings for printing ventilator splitters:

- Layer height: 0.2mm (95% of the stock nozzles on 3D printers are of 0.4mm-keep layer height as 0.5 x nozzle diameter for optimal strength and print time.)

- Perimeters/outer walls: 5 perimeters are suggested for maximum strength.

- Infill: 50-60% suggested for maximum strength. Anything more than this is redundant.

- Path overlap: Max 30%. This setting ensures the walls of the part do not strip off.

- Print speed: 40-60mm/sec is ideal for most printers. The more rigid the frame of a printer is, the higher speed it can print at.

  - Keep the first layer speed as low as 20-30 mm/sec to make sure the first layer sticks

  - Increasing extrusion ratio to 110-120% can also help to stick the first layer to the bed.

- Printing temperature: 190-220 deg Centigrade for most commercial PLA brands.

- No supports needed for both the variants of splitters.

- The splitters can be oriented according to the below images while printing, to prevent cracking.

Please contact

- Mr. Srinivas Kumar,

Ph:  9980911558, Email: Srini@wizardry.co.in

- Mr. Harsha,

Ph: 8016367306, Email: harsha@soch3d.com, harshamadalam@gmail.com

for any further clarifications.